## Slide 1: EVALUATION OF A COMMERCIAL AI BASED SOFTWARE FOR SYNTHETIC CT IMAGE GENERATION
Dr Laura O’Connor Consultant Radiation Therapist Calvary Mater Hospital, Newcastle

## Slide 2: PURPOSE & METHODS
AI based sCT has been shown to be highly accurate in previous studies
Investigate commercially available AI based sCT
Application to pelvis and brain
  40 pelvis (20 prostate, 20 greater pelvis {10m, 10f})
  8 brain (7 primary brain, 1 metastatic)

## Slide 3: METHODS Imaging
Pelvis
Brain

## Slide 4: METHODS Dosimetric Evaluation
Retrospective evaluation
CT gold-standard
  Eclipse planning system
  Accuros calculation model
Assessed:
  %DD at ICRU reference point
    (DsCT-DCT)/DCT * 100%
  Relevant DVH statistics
  3D Gamma
    3%/2mm
    2%/1mm
    1%/1mm

## Slide 5: METHODS Image Guidance Considerations
Auto-registration
DRR key matching points

## Slide 6: METHODS sCT Generation
Deep learning: cGAN
3 outputs:
  120kV equivalent
  Electron density
  Mass density
*Hoesl M, Escobar Corral N & Mistry N. White paper: MR-based Synthetic CT reimagined. Siemens Healthineers

## Slide 7
CT
T1 VIBE DIXON
sCT

## Slide 8: METHODS Brain Bed Model

## Slide 9: RESULTS Pelvis
Dose Difference at ICRU
  AVE = 0.2% (SD 0.6)
  ABS AVE = 0.5% (SD 0.4)
DVH statistics
  AVE = 1.0% (SD 1.0)
3D Gamma
  3%/2mm = 99.98 (SD 0.04)
  2%/1mm = 99.78 (SD 0.24)
  1%/1mm = 96.62 (SD 2.56)

## Slide 10: 3D GAMMA 1%/1mm

## Slide 11: RESULTS Brain
CT
sCT
Dose Difference at ICRU
  AVE = 0.2% (SD 0.5)
  ABS AVE = 0.4% (SD 0.3)
DVH statistics
  AVE = 1.0% (SD 1.8)
3D Gamma
  3%/2mm = 99.79 (SD 0.19)
  2%/1mm = 97.99 (SD 1.53)
  1%/1mm = 92.06 (SD 4.97)

## Slide 12: RESULTS Brain
CT
sCT
Dose Difference at ICRU
  AVE = -0.4% (SD 0.7)
  ABS AVE = 0.7% (SD 0.5)
DVH statistics
  AVE = -0.7% (SD 2.1)
3D Gamma
  3%/2mm = 99.73 (SD 0.18)
  2%/1mm = 97.57 (SD 1.70)
  1%/1mm = 87.30 (SD 5.89

## Slide 13: 3D GAMMA 1%/1mm

## Slide 14: 3D GAMMA 1%/1mm

## Slide 15
RESULTS Brain - Bone

## Slide 16: RESULTS Image Guidance
 | x | y | z | pitch | roll | rot
Pelvis |  |  |  |  |  | 
Mean (mm) | 0.7 | 0.3 | 0.3 | 0.4 | 0.1 | 0.1
SD | 0.4 | 0.3 | 0.3 | 0.4 | 0.1 | 0.2
Brain |  |  |  |  |  | 
Mean (mm) | 0.13 | 0.29 | 0.31 | 0.19 | 0.14 | 0.11
SD | 0.13 | 0.16 | 0.49 | 0.14 | 0.13 | 0.10
 | Pubic sympysis |  |  | LT Sacro Iliac joint | RT Sacro Iliac joint | L5 Ant Vertebral body |  | Sacral/coccyx | Sup Sacrum | 
 | x | y | z | x | x | y | z | y | y | z
Mean (mm) | 1.0 | 1.1 | 1.0 | 1.2 | 1.3 | 1.7 | 2.2 | 1.6 | 2.5 | 1.6
stdev | 2.7 | 2.7 | 2.7 | 1.4 | 1.1 | 2.0 | 3.1 | 2.5 | 3.0 | 1.6

## Slide 17: FUTURE APPLICATIONS
+Hyperarc Brain
+Pituitary and cranio-pharyngioma
01
+Utilize 18-channel ultra-flex coils
02
+Clinical Application
03
+CBCT - Mri image guidance
04

## Slide 18: THANK YOU
Michaela Hoesl
Peter Greer
Claire Monk
Kate Skehan
Siemens Healthineers
Department of Radiation Oncology, Calvary Mater Hospital